## Sheet1
1 | Power supply | Check if all voltages are in range with DMM | Acquire waveforms with oscilloscope | ok for dut1, dut2 with damageg both pmics
2 | Functionality | Check buttons and LEDs functionality | Check working of Over range indication - force ramp on R112/R113 with AD. Check how long is delay, and how long is latch with oscilloscope | needed to change q101 pol;arity and resistor values
2 | DC attenuator | Force 10V differential on input, adjust so bot N and P outputs are equal in reference to GND. Also trim if attenuation is not equal. Use multimeter for probing. Both fow x10 and x100 attenuator |  | ok for now, used 50V input from PSU
3 | AC Attenuator | Perform bode plot with attenuator and amps and ajust frequency response
4 | Gain trim | Check output with DC input |  | hard to do without lab power supply
5 | Offset trim | Trim offset when inputs are shorted, check if same for both x10 and x100
5Mhz filter check??

## Arkusz1
1 | Bode plot of analog discovery
x100 before changes
after fix:
x10 P after fix
x10 N before
x10 N after
x100 N fixed

## Arkusz2
f | M |  | M
Hz | dB |  | dB
1000 | 0.025307 | 1000 | 0.098708 | 0.073402
1010.1781 | 0.025309 | 1010.1781 | 0.096929 | 0.07162
1020.4598 | 0.025528 | 1020.4598 | 0.094557 | 0.06903
1030.8461 | 0.025812 | 1030.8461 | 0.093515 | 0.067703
1041.3381 | 0.02598 | 1041.3381 | 0.092367 | 0.066388
1051.9369 | 0.026276 | 1051.9369 | 0.097623 | 0.071347
1062.6436 | 0.026406 | 1062.6436 | 0.097766 | 0.071359
1073.4593 | 0.026502 | 1073.4593 | 0.096358 | 0.069856
1084.385 | 0.026906 | 1084.385 | 0.090609 | 0.063704
1095.422 | 0.027265 | 1095.422 | 0.096338 | 0.069072
1106.5713 | 0.027498 | 1106.5713 | 0.094541 | 0.067043
1117.8341 | 0.027729 | 1117.8341 | 0.098692 | 0.070963
1129.2115 | 0.027781 | 1129.2115 | 0.096803 | 0.069022
1140.7047 | 0.028265 | 1140.7047 | 0.096946 | 0.068682
1152.3149 | 0.028753 | 1152.3149 | 0.095119 | 0.066366
1164.0432 | 0.029023 | 1164.0432 | 0.097366 | 0.068343
1175.8909 | 0.029374 | 1175.8909 | 0.098157 | 0.068783
1187.8592 | 0.029694 | 1187.8592 | 0.096163 | 0.066468
1199.9494 | 0.029733 | 1199.9494 | 0.093764 | 0.064031
1212.1625 | 0.030255 | 1212.1625 | 0.095278 | 0.065023
1224.5 | 0.03051 | 1224.5 | 0.098266 | 0.067755
1236.9631 | 0.030704 | 1236.9631 | 0.094243 | 0.063539
1249.553 | 0.03125 | 1249.553 | 0.10205 | 0.070802
1262.271 | 0.031624 | 1262.271 | 0.098135 | 0.066511
1275.1185 | 0.031867 | 1275.1185 | 0.091778 | 0.059911
1288.0968 | 0.032055 | 1288.0968 | 0.091087 | 0.059033
1301.2071 | 0.032451 | 1301.2071 | 0.10256 | 0.070111
1314.4509 | 0.032838 | 1314.4509 | 0.090382 | 0.057545
1327.8295 | 0.033312 | 1327.8295 | 0.093489 | 0.060177
1341.3443 | 0.033732 | 1341.3443 | 0.10425 | 0.070513
1354.9966 | 0.034217 | 1354.9966 | 0.094722 | 0.060505
1368.7878 | 0.034317 | 1368.7878 | 0.10682 | 0.072503
1382.7195 | 0.035017 | 1382.7195 | 0.077978 | 0.042961
1396.7929 | 0.035101 | 1396.7929 | 0.089531 | 0.05443
1411.0096 | 0.035591 | 1411.0096 | 0.075394 | 0.039804
1425.3709 | 0.036353 | 1425.3709 | 0.073314 | 0.036961
1439.8785 | 0.036461 | 1439.8785 | 0.1142 | 0.077736
1454.5337 | 0.036915 | 1454.5337 | 0.0901 | 0.053184
1469.338 | 0.036899 | 1469.338 | 0.089581 | 0.052682
1484.2931 | 0.037646 | 1484.2931 | 0.087476 | 0.04983
1499.4003 | 0.038252 | 1499.4003 | 0.090988 | 0.052736
1514.6613 | 0.038386 | 1514.6613 | 0.11405 | 0.075666
1530.0777 | 0.039246 | 1530.0777 | 0.10143 | 0.062183
1545.6509 | 0.039658 | 1545.6509 | 0.082384 | 0.042727
1561.3827 | 0.040033 | 1561.3827 | 0.11452 | 0.074485
1577.2746 | 0.040429 | 1577.2746 | 0.088174 | 0.047745
1593.3282 | 0.040926 | 1593.3282 | 0.080848 | 0.039922
1609.5452 | 0.041404 | 1609.5452 | 0.068737 | 0.027333
1625.9273 | 0.042166 | 1625.9273 | 0.10282 | 0.060656
1642.4761 | 0.042253 | 1642.4761 | 0.099628 | 0.057374
1659.1934 | 0.042763 | 1659.1934 | 0.074352 | 0.031588
1676.0808 | 0.043439 | 1676.0808 | 0.093682 | 0.050243
1693.14 | 0.044277 | 1693.14 | 0.073218 | 0.028941
1710.373 | 0.044993 | 1710.373 | 0.10445 | 0.059452
1727.7813 | 0.045496 | 1727.7813 | 0.10211 | 0.056614
1745.3668 | 0.045275 | 1745.3668 | 0.094549 | 0.049274
1763.1312 | 0.046166 | 1763.1312 | 0.092355 | 0.046189
1781.0765 | 0.046704 | 1781.0765 | 0.11791 | 0.071205
1799.2045 | 0.047339 | 1799.2045 | 0.099727 | 0.052388
1817.5169 | 0.04785 | 1817.5169 | 0.11423 | 0.066377
1836.0157 | 0.048775 | 1836.0157 | 0.090858 | 0.042083
1854.7029 | 0.049362 | 1854.7029 | 0.095968 | 0.046606
1873.5802 | 0.050038 | 1873.5802 | 0.092154 | 0.042116
1892.6496 | 0.050549 | 1892.6496 | 0.10133 | 0.050781
1911.9132 | 0.051293 | 1911.9132 | 0.10396 | 0.052671
1931.3728 | 0.052017 | 1931.3728 | 0.10211 | 0.050093
1951.0304 | 0.05302 | 1951.0304 | 0.096423 | 0.043403
1970.8882 | 0.053592 | 1970.8882 | 0.10337 | 0.049783
1990.948 | 0.054287 | 1990.948 | 0.090691 | 0.036403
2011.212 | 0.054919 | 2011.212 | 0.09885 | 0.043931
2031.6823 | 0.055688 | 2031.6823 | 0.09136 | 0.035672
2052.3609 | 0.056312 | 2052.3609 | 0.097564 | 0.041252
2073.25 | 0.05703 | 2073.25 | 0.094441 | 0.037411
2094.3517 | 0.057878 | 2094.3517 | 0.096651 | 0.038773
2115.6682 | 0.058368 | 2115.6682 | 0.096771 | 0.038403
2137.2016 | 0.059075 | 2137.2016 | 0.095434 | 0.036358
2158.9543 | 0.05988 | 2158.9543 | 0.099559 | 0.039678
2180.9283 | 0.060507 | 2180.9283 | 0.095053 | 0.034546
2203.1259 | 0.061543 | 2203.1259 | 0.096996 | 0.035453
2225.5495 | 0.062275 | 2225.5495 | 0.099071 | 0.036796
2248.2013 | 0.062814 | 2248.2013 | 0.098191 | 0.035377
2271.0837 | 0.063674 | 2271.0837 | 0.09703 | 0.033355
2294.199 | 0.065228 | 2294.199 | 0.095199 | 0.029971
2317.5495 | 0.066043 | 2317.5495 | 0.093388 | 0.027345
2341.1377 | 0.066729 | 2341.1377 | 0.097285 | 0.030556
2364.966 | 0.067497 | 2364.966 | 0.095491 | 0.027994
2389.0368 | 0.068363 | 2389.0368 | 0.10151 | 0.033142
2413.3526 | 0.069122 | 2413.3526 | 0.096187 | 0.027065
2437.9159 | 0.070565 | 2437.9159 | 0.099156 | 0.028591
2462.7292 | 0.071651 | 2462.7292 | 0.10091 | 0.029263
2487.7951 | 0.072107 | 2487.7951 | 0.095708 | 0.023601
2513.116 | 0.073019 | 2513.116 | 0.096153 | 0.023134
2538.6947 | 0.073736 | 2538.6947 | 0.099074 | 0.025338
2564.5338 | 0.075415 | 2564.5338 | 0.11136 | 0.03594
2590.6358 | 0.076292 | 2590.6358 | 0.097705 | 0.021414
2617.0035 | 0.076782 | 2617.0035 | 0.097423 | 0.020641
2643.6396 | 0.077545 | 2643.6396 | 0.10609 | 0.028546
2670.5467 | 0.079361 | 2670.5467 | 0.10638 | 0.027019
2697.7278 | 0.08025 | 2697.7278 | 0.095921 | 0.015671
2725.1854 | 0.080806 | 2725.1854 | 0.10572 | 0.024912
2752.9226 | 0.081811 | 2752.9226 | 0.096753 | 0.014942
2780.9421 | 0.083511 | 2780.9421 | 0.094486 | 0.010975
2809.2467 | 0.084082 | 2809.2467 | 0.091849 | 0.0077678
2837.8394 | 0.085006 | 2837.8394 | 0.094084 | 0.009078
2866.7232 | 0.086981 | 2866.7232 | 0.11859 | 0.03161
2895.9009 | 0.087573 | 2895.9009 | 0.1019 | 0.01433
2925.3756 | 0.088301 | 2925.3756 | 0.075193 | -0.013108
2955.1503 | 0.090177 | 2955.1503 | 0.086894 | -0.0032831
2985.2281 | 0.090999 | 2985.2281 | 0.082918 | -0.0080811
3015.612 | 0.091743 | 3015.612 | 0.11556 | 0.023819
3046.3051 | 0.093418 | 3046.3051 | 0.1037 | 0.010286
3077.3107 | 0.094762 | 3077.3107 | 0.070443 | -0.024319
3108.6318 | 0.095288 | 3108.6318 | 0.079732 | -0.015556
3140.2717 | 0.09698 | 3140.2717 | 0.090913 | -0.0060669
3172.2336 | 0.09776 | 3172.2336 | 0.096579 | -0.0011807
3204.5208 | 0.098541 | 3204.5208 | 0.088305 | -0.010235
3237.1367 | 0.1006 | 3237.1367 | 0.086061 | -0.014534
3270.0845 | 0.10125 | 3270.0845 | 0.10784 | 0.0065858
3303.3677 | 0.10296 | 3303.3677 | 0.096123 | -0.0068411
3336.9897 | 0.10388 | 3336.9897 | 0.091603 | -0.012276
3370.9538 | 0.10463 | 3370.9538 | 0.099361 | -0.0052653
3405.2636 | 0.1066 | 3405.2636 | 0.092275 | -0.014328
3439.9227 | 0.1073 | 3439.9227 | 0.090937 | -0.016364
3474.9345 | 0.10907 | 3474.9345 | 0.078877 | -0.030194
3510.3026 | 0.10972 | 3510.3026 | 0.099277 | -0.010445
3546.0308 | 0.11095 | 3546.0308 | 0.096145 | -0.014806
3582.1226 | 0.11245 | 3582.1226 | 0.092624 | -0.019822
3618.5817 | 0.11304 | 3618.5817 | 0.10469 | -0.0083529
3655.4119 | 0.11491 | 3655.4119 | 0.10572 | -0.0091912
3692.617 | 0.11558 | 3692.617 | 0.096062 | -0.019516
3730.2007 | 0.11753 | 3730.2007 | 0.0966 | -0.020926
3768.167 | 0.11827 | 3768.167 | 0.097902 | -0.020366
3806.5197 | 0.12021 | 3806.5197 | 0.098253 | -0.021955
3845.2628 | 0.12094 | 3845.2628 | 0.098937 | -0.022004
3884.4001 | 0.12295 | 3884.4001 | 0.097995 | -0.024955
3923.9359 | 0.12362 | 3923.9359 | 0.10117 | -0.022453
3963.874 | 0.12508 | 3963.874 | 0.096665 | -0.028414
4004.2186 | 0.12598 | 4004.2186 | 0.09304 | -0.032936
4044.9739 | 0.12774 | 4044.9739 | 0.1009 | -0.026842
4086.1439 | 0.12846 | 4086.1439 | 0.098012 | -0.030448
4127.733 | 0.13039 | 4127.733 | 0.10091 | -0.02948
4169.7454 | 0.13088 | 4169.7454 | 0.09284 | -0.038038
4212.1854 | 0.13281 | 4212.1854 | 0.094796 | -0.038012
4255.0574 | 0.13333 | 4255.0574 | 0.096559 | -0.03677
4298.3657 | 0.1353 | 4298.3657 | 0.098216 | -0.037086
4342.1148 | 0.13581 | 4342.1148 | 0.095825 | -0.039984
4386.3092 | 0.13767 | 4386.3092 | 0.096416 | -0.041256
4430.9534 | 0.1372 | 4430.9534 | 0.096258 | -0.040938
4476.052 | 0.14003 | 4476.052 | 0.097818 | -0.042212
4521.6096 | 0.14092 | 4521.6096 | 0.097705 | -0.043219
4567.6309 | 0.14214 | 4567.6309 | 0.092362 | -0.049783
4614.1206 | 0.14376 | 4614.1206 | 0.097728 | -0.04603
4661.0835 | 0.14467 | 4661.0835 | 0.096477 | -0.048189
4708.5244 | 0.14649 | 4708.5244 | 0.099788 | -0.046705
4756.4481 | 0.14429 | 4756.4481 | 0.092822 | -0.051469
4804.8596 | 0.1488 | 4804.8596 | 0.094753 | -0.054043
4853.7638 | 0.14909 | 4853.7638 | 0.095123 | -0.053967
4903.1658 | 0.15035 | 4903.1658 | 0.094885 | -0.055465
4953.0706 | 0.152 | 4953.0706 | 0.095135 | -0.056866
5003.4834 | 0.14825 | 5003.4834 | 0.096397 | -0.051853
5054.4092 | 0.15465 | 5054.4092 | 0.096271 | -0.058384
5105.8534 | 0.15433 | 5105.8534 | 0.09593 | -0.058398
5157.8212 | 0.1559 | 5157.8212 | 0.096493 | -0.059407
5210.3179 | 0.15787 | 5210.3179 | 0.091756 | -0.066117
5263.3489 | 0.15464 | 5263.3489 | 0.09161 | -0.06303
5316.9197 | 0.15971 | 5316.9197 | 0.093547 | -0.066164
5371.0357 | 0.16062 | 5371.0357 | 0.093076 | -0.06754
5425.7026 | 0.16122 | 5425.7026 | 0.094354 | -0.066862
5480.9258 | 0.16291 | 5480.9258 | 0.10462 | -0.058288
5536.7111 | 0.16239 | 5536.7111 | 0.096469 | -0.065922
5593.0642 | 0.16443 | 5593.0642 | 0.096209 | -0.068223
5649.9908 | 0.16595 | 5649.9908 | 0.094838 | -0.07111
5707.4969 | 0.16373 | 5707.4969 | 0.090719 | -0.073014
5765.5882 | 0.16724 | 5765.5882 | 0.095699 | -0.071543
5824.2708 | 0.16878 | 5824.2708 | 0.0931 | -0.075677
5883.5507 | 0.16502 | 5883.5507 | 0.093395 | -0.071624
5943.434 | 0.16976 | 5943.434 | 0.09283 | -0.07693
6003.9267 | 0.17126 | 6003.9267 | 0.087959 | -0.083297
6065.0351 | 0.16702 | 6065.0351 | 0.091419 | -0.075596
6126.7655 | 0.17246 | 6126.7655 | 0.097945 | -0.074519
6189.1243 | 0.17368 | 6189.1243 | 0.098358 | -0.07532
6252.1176 | 0.16975 | 6252.1176 | 0.0869 | -0.082848
6315.7522 | 0.17437 | 6315.7522 | 0.092545 | -0.081826
6380.0344 | 0.17575 | 6380.0344 | 0.091695 | -0.084056
6444.9709 | 0.17297 | 6444.9709 | 0.083842 | -0.089124
6510.5683 | 0.17648 | 6510.5683 | 0.09159 | -0.084893
6576.8334 | 0.17762 | 6576.8334 | 0.086036 | -0.091586
6643.7729 | 0.17668 | 6643.7729 | 0.10356 | -0.073124
6711.3938 | 0.16344 | 6711.3938 | 0.070645 | -0.092798
6779.7029 | 0.17889 | 6779.7029 | 0.094682 | -0.084209
6848.7072 | 0.17926 | 6848.7072 | 0.09429 | -0.084967
6918.4139 | 0.17179 | 6918.4139 | 0.084406 | -0.087382
6988.8301 | 0.17989 | 6988.8301 | 0.093795 | -0.086095
7059.9629 | 0.18101 | 7059.9629 | 0.096594 | -0.084413
7131.8198 | 0.17882 | 7131.8198 | 0.083903 | -0.094913
7204.408 | 0.16918 | 7204.408 | 0.069886 | -0.099291
7277.7351 | 0.18195 | 7277.7351 | 0.092205 | -0.089745
7351.8084 | 0.18246 | 7351.8084 | 0.092411 | -0.09005
7426.6357 | 0.17586 | 7426.6357 | 0.084787 | -0.091072
7502.2246 | 0.17773 | 7502.2246 | 0.091336 | -0.086392
7578.5828 | 0.18332 | 7578.5828 | 0.092658 | -0.090659
7655.7182 | 0.18282 | 7655.7182 | 0.080679 | -0.10214
7733.6388 | 0.17145 | 7733.6388 | 0.078142 | -0.09331
7812.3523 | 0.1812 | 7812.3523 | 0.085673 | -0.09553
7891.8671 | 0.18428 | 7891.8671 | 0.092789 | -0.091493
7972.1911 | 0.1831 | 7972.1911 | 0.079201 | -0.1039
8053.3327 | 0.16943 | 8053.3327 | 0.073845 | -0.095585
8135.3002 | 0.18292 | 8135.3002 | 0.082135 | -0.10079
8218.1019 | 0.18504 | 8218.1019 | 0.089889 | -0.095155
8301.7464 | 0.18349 | 8301.7464 | 0.089193 | -0.094299
8386.2422 | 0.1703 | 8386.2422 | 0.078806 | -0.091494
8471.5981 | 0.18261 | 8471.5981 | 0.080818 | -0.10179
8557.8226 | 0.18475 | 8557.8226 | 0.084409 | -0.10034
8644.9248 | 0.18439 | 8644.9248 | 0.081697 | -0.10269
8732.9136 | 0.17463 | 8732.9136 | 0.080573 | -0.094057
8821.7979 | 0.18083 | 8821.7979 | 0.081936 | -0.098896
8911.5868 | 0.18403 | 8911.5868 | 0.086648 | -0.097385
9002.2896 | 0.18434 | 9002.2896 | 0.087666 | -0.096678
9093.9157 | 0.17955 | 9093.9157 | 0.08057 | -0.098981
9186.4742 | 0.17465 | 9186.4742 | 0.0796 | -0.095052
9279.9749 | 0.1833 | 9279.9749 | 0.088311 | -0.09499
9374.4272 | 0.18354 | 9374.4272 | 0.082619 | -0.10092
9469.8409 | 0.18308 | 9469.8409 | 0.081 | -0.10208
9566.2257 | 0.17263 | 9566.2257 | 0.069235 | -0.10339
9663.5915 | 0.17914 | 9663.5915 | 0.079178 | -0.099961
9761.9483 | 0.18254 | 9761.9483 | 0.079918 | -0.10263
9861.3062 | 0.18187 | 9861.3062 | 0.092091 | -0.089779
9961.6753 | 0.18025 | 9961.6753 | 0.086434 | -0.093817
10063.066 | 0.16713 | 10063.066 | 0.068687 | -0.098445
10165.4887 | 0.17888 | 10165.4887 | 0.075177 | -0.1037
10268.9538 | 0.18073 | 10268.9538 | 0.08505 | -0.095684
10373.4721 | 0.18009 | 10373.4721 | 0.085036 | -0.09505
10479.0541 | 0.17915 | 10479.0541 | 0.080439 | -0.098715
10585.7107 | 0.16665 | 10585.7107 | 0.063033 | -0.10362
10693.4529 | 0.17652 | 10693.4529 | 0.074131 | -0.10239
10802.2917 | 0.17873 | 10802.2917 | 0.073301 | -0.10543
10912.2383 | 0.17775 | 10912.2383 | 0.073949 | -0.1038
11023.3039 | 0.17732 | 11023.3039 | 0.079635 | -0.097685
11135.5 | 0.16994 | 11135.5 | 0.069092 | -0.10085
11248.838 | 0.16997 | 11248.838 | 0.075828 | -0.094139
11363.3296 | 0.17534 | 11363.3296 | 0.080627 | -0.094708
11478.9864 | 0.1746 | 11478.9864 | 0.077298 | -0.097305
11595.8204 | 0.17452 | 11595.8204 | 0.075094 | -0.099426
11713.8436 | 0.17266 | 11713.8436 | 0.072458 | -0.1002
11833.0681 | 0.16016 | 11833.0681 | 0.063553 | -0.096611
11953.506 | 0.17041 | 11953.506 | 0.073134 | -0.097278
12075.1697 | 0.17204 | 12075.1697 | 0.075043 | -0.096995
12198.0717 | 0.17092 | 12198.0717 | 0.068054 | -0.10287
12322.2246 | 0.1707 | 12322.2246 | 0.074703 | -0.095995
12447.6412 | 0.16917 | 12447.6412 | 0.069204 | -0.099967
12574.3343 | 0.15614 | 12574.3343 | 0.054144 | -0.10199
12702.3169 | 0.16654 | 12702.3169 | 0.061784 | -0.10475
12831.602 | 0.16826 | 12831.602 | 0.067255 | -0.10101
12962.2031 | 0.16684 | 12962.2031 | 0.068829 | -0.098008
13094.1334 | 0.16652 | 13094.1334 | 0.074941 | -0.09158
13227.4066 | 0.16611 | 13227.4066 | 0.074088 | -0.092021
13362.0361 | 0.15888 | 13362.0361 | 0.05729 | -0.10159
13498.036 | 0.15873 | 13498.036 | 0.058989 | -0.099742
13635.4201 | 0.16416 | 13635.4201 | 0.070515 | -0.093647
13774.2025 | 0.16314 | 13774.2025 | 0.051558 | -0.11158
13914.3974 | 0.16248 | 13914.3974 | 0.075047 | -0.08743
14056.0192 | 0.16173 | 14056.0192 | 0.056196 | -0.10554
14199.0825 | 0.16065 | 14199.0825 | 0.075831 | -0.084821
14343.6019 | 0.15073 | 14343.6019 | 0.05426 | -0.096466
14489.5922 | 0.15597 | 14489.5922 | 0.068622 | -0.087348
14637.0684 | 0.15919 | 14637.0684 | 0.043924 | -0.11527
14786.0456 | 0.15769 | 14786.0456 | 0.072882 | -0.084808
14936.5392 | 0.15736 | 14936.5392 | 0.058961 | -0.098403
15088.5645 | 0.1566 | 15088.5645 | 0.066362 | -0.090238
15242.1371 | 0.15635 | 15242.1371 | 0.05796 | -0.098389
15397.2727 | 0.15083 | 15397.2727 | 0.061224 | -0.089608
15553.9874 | 0.14636 | 15553.9874 | 0.037997 | -0.10837
15712.2971 | 0.15418 | 15712.2971 | 0.057058 | -0.097123
15872.2181 | 0.15325 | 15872.2181 | 0.05676 | -0.096489
16033.7668 | 0.1523 | 16033.7668 | 0.046359 | -0.10594
16196.9597 | 0.15213 | 16196.9597 | 0.056537 | -0.095598
16361.8137 | 0.15113 | 16361.8137 | 0.049236 | -0.10189
16528.3455 | 0.15072 | 16528.3455 | 0.050475 | -0.10025
16696.5723 | 0.14365 | 16696.5723 | 0.034195 | -0.10946
16866.5113 | 0.14247 | 16866.5113 | 0.044794 | -0.097673
17038.18 | 0.14889 | 17038.18 | 0.053305 | -0.095581
17211.596 | 0.14817 | 17211.596 | 0.057194 | -0.09098
17386.7769 | 0.14701 | 17386.7769 | 0.04809 | -0.098921
17563.7409 | 0.14657 | 17563.7409 | 0.046911 | -0.099657
17742.5061 | 0.13712 | 17742.5061 | 0.050936 | -0.086187
17923.0907 | 0.14524 | 17923.0907 | 0.045857 | -0.099387
18105.5133 | 0.14382 | 18105.5133 | 0.045568 | -0.098249
18289.7927 | 0.1306 | 18289.7927 | 0.027715 | -0.10289
18475.9476 | 0.14089 | 18475.9476 | 0.038786 | -0.10211
18663.9973 | 0.14299 | 18663.9973 | 0.038227 | -0.10476
18853.9609 | 0.14186 | 18853.9609 | 0.05496 | -0.086903
19045.858 | 0.14168 | 19045.858 | 0.047588 | -0.094093
19239.7083 | 0.13927 | 19239.7083 | 0.056471 | -0.082798
19435.5315 | 0.12657 | 19435.5315 | 0.033929 | -0.092639
19633.3479 | 0.1394 | 19633.3479 | 0.045521 | -0.093878
19833.1777 | 0.13906 | 19833.1777 | 0.031196 | -0.10787
20035.0413 | 0.13267 | 20035.0413 | 0.04367 | -0.089
20238.9596 | 0.13147 | 20238.9596 | 0.032854 | -0.098617
20444.9533 | 0.1378 | 20444.9533 | 0.042857 | -0.094941
20653.0436 | 0.13636 | 20653.0436 | 0.041524 | -0.094841
20863.2519 | 0.13605 | 20863.2519 | 0.043675 | -0.092378
21075.5998 | 0.13593 | 21075.5998 | 0.041641 | -0.09429
21290.1089 | 0.12699 | 21290.1089 | 0.026231 | -0.10076
21506.8013 | 0.1314 | 21506.8013 | 0.035452 | -0.095946
21725.6992 | 0.13389 | 21725.6992 | 0.044255 | -0.089632
21946.8251 | 0.13452 | 21946.8251 | 0.035534 | -0.09899
22170.2016 | 0.13144 | 22170.2016 | 0.028273 | -0.10317
22395.8517 | 0.11831 | 22395.8517 | 0.035979 | -0.082336
22623.7984 | 0.13161 | 22623.7984 | 0.035548 | -0.096062
22854.0652 | 0.13238 | 22854.0652 | 0.039923 | -0.092461
23086.6757 | 0.13068 | 23086.6757 | 0.05101 | -0.079673
23321.6537 | 0.13129 | 23321.6537 | 0.032695 | -0.098591
23559.0234 | 0.12767 | 23559.0234 | 0.018461 | -0.10921
23798.809 | 0.12002 | 23798.809 | 0.026057 | -0.093965
24041.0351 | 0.12908 | 24041.0351 | 0.043726 | -0.085354
24285.7267 | 0.12938 | 24285.7267 | 0.035917 | -0.093459
24532.9087 | 0.12884 | 24532.9087 | 0.039508 | -0.089331
24782.6066 | 0.12823 | 24782.6066 | 0.029083 | -0.09915
25034.8459 | 0.12239 | 25034.8459 | 0.025747 | -0.096645
25289.6526 | 0.12071 | 25289.6526 | 0.034956 | -0.085759
25547.0526 | 0.12746 | 25547.0526 | 0.034217 | -0.09324
25807.0726 | 0.12693 | 25807.0726 | 0.043668 | -0.083261
26069.739 | 0.12677 | 26069.739 | 0.033333 | -0.093437
26335.0788 | 0.12671 | 26335.0788 | 0.029926 | -0.096783
26603.1193 | 0.11768 | 26603.1193 | 0.021008 | -0.096667
26873.888 | 0.12223 | 26873.888 | 0.020632 | -0.1016
27147.4126 | 0.12589 | 27147.4126 | 0.018653 | -0.10723
27423.7211 | 0.12505 | 27423.7211 | 0.024303 | -0.10075
27702.8419 | 0.12518 | 27702.8419 | 0.01795 | -0.10723
27984.8036 | 0.12467 | 27984.8036 | 0.04841 | -0.076257
28269.6351 | 0.12438 | 28269.6351 | 0.021895 | -0.10248
28557.3657 | 0.12047 | 28557.3657 | 0.01668 | -0.10379
28848.0248 | 0.11302 | 28848.0248 | -0.0018951 | -0.11492
29141.6423 | 0.12304 | 29141.6423 | 0.00023333 | -0.1228
29438.2482 | 0.123 | 29438.2482 | 0.026047 | -0.09695
29737.873 | 0.12232 | 29737.873 | 0.00704 | -0.11528
30040.5475 | 0.12285 | 30040.5475 | 0.011031 | -0.11182
30346.3025 | 0.11687 | 30346.3025 | 0.019841 | -0.097028
30655.1696 | 0.11545 | 30655.1696 | 0.0085844 | -0.10687
30967.1803 | 0.12216 | 30967.1803 | 0.019565 | -0.1026
31282.3667 | 0.12129 | 31282.3667 | 0.016239 | -0.10505
31600.7611 | 0.12171 | 31600.7611 | 0.013467 | -0.10824
31922.3962 | 0.12088 | 31922.3962 | 0.025976 | -0.094902
32247.3048 | 0.10627 | 32247.3048 | -0.0011232 | -0.1074
32575.5204 | 0.12088 | 32575.5204 | 0.017572 | -0.1033
32907.0767 | 0.12171 | 32907.0767 | 0.009296 | -0.11241
33242.0075 | 0.11953 | 33242.0075 | 0.045813 | -0.073712
33580.3473 | 0.10524 | 33580.3473 | 0.010907 | -0.094335
33922.1307 | 0.11982 | 33922.1307 | 0.020247 | -0.099568
34267.3928 | 0.12046 | 34267.3928 | 0.036425 | -0.084038
34616.169 | 0.11975 | 34616.169 | 0.009654 | -0.1101
34968.4951 | 0.12111 | 34968.4951 | 0.012513 | -0.10859
35324.4073 | 0.11728 | 35324.4073 | 0.012356 | -0.10493
35683.9419 | 0.10953 | 35683.9419 | 0.026432 | -0.083094
36047.1358 | 0.11937 | 36047.1358 | 0.016668 | -0.1027
36414.0264 | 0.12047 | 36414.0264 | 0.0193 | -0.10117
36784.6513 | 0.11987 | 36784.6513 | 0.014559 | -0.10531
37159.0483 | 0.12068 | 37159.0483 | 0.031293 | -0.08939
37537.2561 | 0.11051 | 37537.2561 | 0.012609 | -0.097906
37919.3132 | 0.1174 | 37919.3132 | 0.018722 | -0.098676
38305.259 | 0.12008 | 38305.259 | 0.023394 | -0.096682
38695.1329 | 0.11978 | 38695.1329 | 0.020949 | -0.098832
39088.9751 | 0.11961 | 39088.9751 | 0.037012 | -0.082597
39486.8257 | 0.12015 | 39486.8257 | 0.029585 | -0.090563
39888.7258 | 0.1182 | 39888.7258 | 0.0055873 | -0.11261
40294.7164 | 0.10517 | 40294.7164 | 0.017671 | -0.087499
40704.8392 | 0.11935 | 40704.8392 | 0.017957 | -0.1014
41119.1362 | 0.11992 | 41119.1362 | 0.032192 | -0.087727
41537.6501 | 0.12025 | 41537.6501 | 0.037419 | -0.082828
41960.4235 | 0.12023 | 41960.4235 | 0.0042402 | -0.11599
42387.5 | 0.11781 | 42387.5 | 0.035361 | -0.082448
42818.9234 | 0.10935 | 42818.9234 | 0.024635 | -0.084717
43254.7377 | 0.12028 | 43254.7377 | 0.032043 | -0.088241
43694.9879 | 0.12062 | 43694.9879 | 0.021618 | -0.098998
44139.7189 | 0.11987 | 44139.7189 | 0.028969 | -0.090897
44588.9765 | 0.12071 | 44588.9765 | 0.022925 | -0.097784
45042.8066 | 0.11106 | 45042.8066 | 0.012488 | -0.098573
45501.2558 | 0.11837 | 45501.2558 | 0.015267 | -0.1031
45964.3712 | 0.12161 | 45964.3712 | 0.02335 | -0.098258
46432.2002 | 0.12089 | 46432.2002 | 0.019629 | -0.10126
46904.7908 | 0.1221 | 46904.7908 | 0.022369 | -0.099735
47382.1915 | 0.11611 | 47382.1915 | 0.029 | -0.087107
47864.4511 | 0.1169 | 47864.4511 | 0.014772 | -0.10213
48351.6193 | 0.12217 | 48351.6193 | 0.01823 | -0.10394
48843.7459 | 0.12206 | 48843.7459 | 0.026576 | -0.095488
49340.8814 | 0.1222 | 49340.8814 | 0.019735 | -0.10246
49843.0768 | 0.12119 | 49843.0768 | 0.019817 | -0.10137
50350.3836 | 0.10623 | 50350.3836 | 0.024345 | -0.081887
50862.8537 | 0.12097 | 50862.8537 | 0.044516 | -0.076458
51380.5399 | 0.12261 | 51380.5399 | 0.039563 | -0.083042
51903.4951 | 0.12263 | 51903.4951 | 0.054515 | -0.068119
52431.773 | 0.1233 | 52431.773 | 0.040699 | -0.082605
52965.4277 | 0.12142 | 52965.4277 | 0.011052 | -0.11037
53504.514 | 0.11227 | 53504.514 | 0.009112 | -0.10316
54049.0871 | 0.12272 | 54049.0871 | 0.046428 | -0.076289
54599.203 | 0.12461 | 54599.203 | 0.022629 | -0.10198
55154.918 | 0.12347 | 55154.918 | 0.033487 | -0.089986
55716.2891 | 0.12512 | 55716.2891 | 0.041618 | -0.083498
56283.3739 | 0.1159 | 56283.3739 | 0.039457 | -0.076448
56856.2305 | 0.12288 | 56856.2305 | 0.037507 | -0.085372
57434.9177 | 0.12573 | 57434.9177 | 0.032612 | -0.093116
58019.4949 | 0.12528 | 58019.4949 | 0.043322 | -0.081958
58610.0219 | 0.12578 | 58610.0219 | 0.039572 | -0.086204
59206.5593 | 0.12109 | 59206.5593 | 0.035439 | -0.085654
59809.1683 | 0.12114 | 59809.1683 | 0.029495 | -0.091649
60417.9108 | 0.12684 | 60417.9108 | 0.036826 | -0.090017
61032.849 | 0.12625 | 61032.849 | 0.048842 | -0.077412
61654.0462 | 0.12771 | 61654.0462 | 0.040151 | -0.087559
62281.566 | 0.12114 | 62281.566 | 0.03229 | -0.088847
62915.4726 | 0.12264 | 62915.4726 | 0.041125 | -0.081511
63555.8313 | 0.12857 | 63555.8313 | 0.040598 | -0.087977
64202.7076 | 0.12871 | 64202.7076 | 0.044626 | -0.084084
64856.1678 | 0.1288 | 64856.1678 | 0.040787 | -0.088011
65516.279 | 0.1292 | 65516.279 | 0.044737 | -0.084464
66183.1088 | 0.12937 | 66183.1088 | 0.03641 | -0.092962
66856.7258 | 0.12231 | 66856.7258 | 0.043325 | -0.078989
67537.1988 | 0.12582 | 67537.1988 | 0.030759 | -0.095061
68224.5977 | 0.13112 | 68224.5977 | 0.04643 | -0.084685
68918.9931 | 0.13059 | 68918.9931 | 0.046193 | -0.084394
69620.456 | 0.13149 | 69620.456 | 0.049875 | -0.081614
70329.0585 | 0.13198 | 70329.0585 | 0.061403 | -0.07058
71044.8732 | 0.12185 | 71044.8732 | 0.046161 | -0.075687
71767.9736 | 0.1294 | 71767.9736 | 0.042999 | -0.086404
72498.4337 | 0.13281 | 72498.4337 | 0.052076 | -0.080732
73236.3284 | 0.13336 | 73236.3284 | 0.054834 | -0.07853
73981.7336 | 0.13361 | 73981.7336 | 0.050408 | -0.083207
74734.7255 | 0.13053 | 74734.7255 | 0.043574 | -0.086956
75495.3814 | 0.12457 | 75495.3814 | 0.046396 | -0.078175
76263.7794 | 0.13399 | 76263.7794 | 0.0742 | -0.059786
77039.9982 | 0.13457 | 77039.9982 | 0.061128 | -0.073446
77824.1173 | 0.13447 | 77824.1173 | 0.052697 | -0.081769
78616.2173 | 0.13427 | 78616.2173 | 0.068491 | -0.06578
79416.3794 | 0.12236 | 79416.3794 | 0.038414 | -0.083942
80224.6856 | 0.13546 | 80224.6856 | 0.07617 | -0.059291
81041.2188 | 0.13635 | 81041.2188 | 0.074578 | -0.06177
81866.0627 | 0.13623 | 81866.0627 | 0.063783 | -0.072442
82699.302 | 0.13583 | 82699.302 | 0.043765 | -0.092066
83541.022 | 0.12521 | 83541.022 | 0.049617 | -0.075597
84391.3091 | 0.1382 | 84391.3091 | 0.065385 | -0.072817
85250.2505 | 0.13758 | 85250.2505 | 0.056969 | -0.080611
86117.9343 | 0.13852 | 86117.9343 | 0.061194 | -0.077328
86994.4494 | 0.13475 | 86994.4494 | 0.072439 | -0.062313
87879.8858 | 0.13439 | 87879.8858 | 0.070437 | -0.063951
88774.3342 | 0.1393 | 88774.3342 | 0.072829 | -0.066467
89677.8864 | 0.13883 | 89677.8864 | 0.058562 | -0.080273
90590.635 | 0.13968 | 90590.635 | 0.063066 | -0.076618
91512.6736 | 0.12472 | 91512.6736 | 0.057197 | -0.067522
92444.0968 | 0.14037 | 92444.0968 | 0.070761 | -0.069606
93385.0002 | 0.14095 | 93385.0002 | 0.06688 | -0.074068
94335.4801 | 0.14133 | 94335.4801 | 0.085878 | -0.055448
95295.634 | 0.13499 | 95295.634 | 0.070462 | -0.064531
96265.5605 | 0.14031 | 96265.5605 | 0.067908 | -0.0724
97245.359 | 0.14088 | 97245.359 | 0.087224 | -0.053657
98235.1299 | 0.14219 | 98235.1299 | 0.10424 | -0.037949
99234.9748 | 0.14336 | 99234.9748 | 0.079522 | -0.063835
100244.9963 | 0.13763 | 100244.9963 | 0.082571 | -0.055057
101265.2977 | 0.13765 | 101265.2977 | 0.071699 | -0.065947
102295.9839 | 0.14414 | 102295.9839 | 0.098783 | -0.04536
103337.1605 | 0.14336 | 103337.1605 | 0.096558 | -0.0468
104388.9343 | 0.14357 | 104388.9343 | 0.084412 | -0.059157
105451.4131 | 0.14329 | 105451.4131 | 0.097955 | -0.045336
106524.7059 | 0.13228 | 106524.7059 | 0.065083 | -0.067201
107608.9228 | 0.13982 | 107608.9228 | 0.083328 | -0.056495
108704.1749 | 0.14403 | 108704.1749 | 0.077535 | -0.066495
109810.5746 | 0.14537 | 109810.5746 | 0.088479 | -0.056887
110928.2353 | 0.14576 | 110928.2353 | 0.094786 | -0.050974
112057.2716 | 0.14268 | 112057.2716 | 0.10022 | -0.04246
113197.7994 | 0.13548 | 113197.7994 | 0.089007 | -0.046469
114349.9355 | 0.14626 | 114349.9355 | 0.099645 | -0.046614
115513.7982 | 0.14549 | 115513.7982 | 0.092994 | -0.052495
116689.5068 | 0.14609 | 116689.5068 | 0.11594 | -0.030145
117877.1818 | 0.14474 | 117877.1818 | 0.094362 | -0.050376
119076.945 | 0.13176 | 119076.945 | 0.088088 | -0.043675
120288.9196 | 0.14643 | 120288.9196 | 0.096212 | -0.050214
121513.2297 | 0.14529 | 121513.2297 | 0.10406 | -0.041227
122750.0009 | 0.14792 | 122750.0009 | 0.10478 | -0.043136
123999.3601 | 0.14599 | 123999.3601 | 0.097383 | -0.048603
125261.4354 | 0.13459 | 125261.4354 | 0.091627 | -0.042965
126536.3562 | 0.14755 | 126536.3562 | 0.10084 | -0.04671
127824.2532 | 0.147 | 127824.2532 | 0.11079 | -0.036211
129125.2586 | 0.14695 | 129125.2586 | 0.11154 | -0.035406
130439.5056 | 0.14355 | 130439.5056 | 0.11902 | -0.024525
131767.1292 | 0.13951 | 131767.1292 | 0.10438 | -0.035133
133108.2655 | 0.14622 | 133108.2655 | 0.12057 | -0.025652
134463.0519 | 0.14577 | 134463.0519 | 0.11934 | -0.026429
135831.6275 | 0.14698 | 135831.6275 | 0.12372 | -0.023263
137214.1325 | 0.13294 | 137214.1325 | 0.088888 | -0.044055
138610.7088 | 0.14612 | 138610.7088 | 0.12083 | -0.025288
140021.4995 | 0.14621 | 140021.4995 | 0.11246 | -0.033749
141446.6494 | 0.14599 | 141446.6494 | 0.12203 | -0.02396
142886.3045 | 0.13989 | 142886.3045 | 0.10409 | -0.0358
144340.6126 | 0.14252 | 144340.6126 | 0.12108 | -0.02144
145809.7228 | 0.14435 | 145809.7228 | 0.10957 | -0.034779
147293.7857 | 0.14581 | 147293.7857 | 0.12118 | -0.024627
148792.9534 | 0.14127 | 148792.9534 | 0.1254 | -0.015876
150307.3799 | 0.14174 | 150307.3799 | 0.1213 | -0.020443
151837.2202 | 0.14509 | 151837.2202 | 0.11868 | -0.02641
153382.6315 | 0.14396 | 153382.6315 | 0.10858 | -0.035381
154943.772 | 0.14098 | 154943.772 | 0.10536 | -0.035624
156520.8019 | 0.14189 | 156520.8019 | 0.11972 | -0.022163
158113.883 | 0.14412 | 158113.883 | 0.13471 | -0.0094118
159723.1786 | 0.14442 | 159723.1786 | 0.12558 | -0.018833
161348.8537 | 0.1352 | 161348.8537 | 0.11322 | -0.021987
162991.0751 | 0.14299 | 162991.0751 | 0.12624 | -0.016752
164650.0111 | 0.14293 | 164650.0111 | 0.14009 | -0.002838
166325.8319 | 0.14277 | 166325.8319 | 0.11481 | -0.027964
168018.7094 | 0.1294 | 168018.7094 | 0.11963 | -0.0097622
169728.8171 | 0.14261 | 169728.8171 | 0.1265 | -0.016103
171456.3304 | 0.14141 | 171456.3304 | 0.13195 | -0.0094619
173201.4264 | 0.13641 | 173201.4264 | 0.13123 | -0.0051803
174964.2842 | 0.13823 | 174964.2842 | 0.12982 | -0.0084122
176745.0845 | 0.14066 | 176745.0845 | 0.13485 | -0.0058069
178544.0099 | 0.1413 | 178544.0099 | 0.1197 | -0.021595
180361.245 | 0.1303 | 180361.245 | 0.11207 | -0.018228
182196.976 | 0.13941 | 182196.976 | 0.13634 | -0.003068
184051.3912 | 0.13932 | 184051.3912 | 0.12357 | -0.015754
185924.6808 | 0.12615 | 185924.6808 | 0.12788 | 0.0017305
187817.0368 | 0.13879 | 187817.0368 | 0.14338 | 0.0045846
189728.6535 | 0.13645 | 189728.6535 | 0.13045 | -0.0059946
191659.7267 | 0.13151 | 191659.7267 | 0.13702 | 0.0055119
193610.4545 | 0.13464 | 193610.4545 | 0.12539 | -0.0092557
195581.037 | 0.13357 | 195581.037 | 0.10856 | -0.02501
197571.6762 | 0.13148 | 197571.6762 | 0.13173 | 0.00025565
199582.5764 | 0.13068 | 199582.5764 | 0.098974 | -0.031703
201613.9436 | 0.11902 | 201613.9436 | 0.093221 | -0.025795
203665.9862 | 0.12895 | 203665.9862 | 0.12344 | -0.0055126
205738.9147 | 0.12977 | 205738.9147 | 0.11667 | -0.013102
207832.9416 | 0.13077 | 207832.9416 | 0.11713 | -0.013644
209948.2817 | 0.12996 | 209948.2817 | 0.12345 | -0.0065024
212085.1519 | 0.12571 | 212085.1519 | 0.11074 | -0.014972
214243.7713 | 0.11791 | 214243.7713 | 0.12923 | 0.011324
216424.3613 | 0.12865 | 216424.3613 | 0.12409 | -0.004555
218627.1456 | 0.1279 | 218627.1456 | 0.13124 | 0.0033452
220852.3499 | 0.12566 | 220852.3499 | 0.12638 | 0.00072152
223100.2026 | 0.12319 | 223100.2026 | 0.13573 | 0.012539
225370.9341 | 0.11315 | 225370.9341 | 0.10623 | -0.0069209
227664.7772 | 0.12145 | 227664.7772 | 0.11747 | -0.0039733
229981.9673 | 0.12227 | 229981.9673 | 0.12904 | 0.0067636
232322.7419 | 0.122 | 232322.7419 | 0.11958 | -0.0024234
234687.3411 | 0.12038 | 234687.3411 | 0.13914 | 0.018764
237076.0074 | 0.11334 | 237076.0074 | 0.12077 | 0.0074318
239488.9857 | 0.11479 | 239488.9857 | 0.13159 | 0.016792
241926.5235 | 0.11903 | 241926.5235 | 0.13898 | 0.019951
244388.8708 | 0.11684 | 244388.8708 | 0.11495 | -0.0018813
246876.28 | 0.11559 | 246876.28 | 0.13932 | 0.023727
249389.0063 | 0.10788 | 249389.0063 | 0.10335 | -0.0045319
251927.3073 | 0.11313 | 251927.3073 | 0.13748 | 0.024354
254491.4433 | 0.11454 | 254491.4433 | 0.10861 | -0.0059344
257081.6773 | 0.11202 | 257081.6773 | 0.11503 | 0.003015
259698.2749 | 0.11112 | 259698.2749 | 0.12863 | 0.01751
262341.5045 | 0.096081 | 262341.5045 | 0.11104 | 0.014956
265011.637 | 0.11018 | 265011.637 | 0.12345 | 0.013271
267708.9464 | 0.10757 | 267708.9464 | 0.11578 | 0.0082172
270433.7092 | 0.10747 | 270433.7092 | 0.11898 | 0.011508
273186.2048 | 0.10427 | 273186.2048 | 0.12443 | 0.020159
275966.7156 | 0.10241 | 275966.7156 | 0.11426 | 0.011853
278775.5266 | 0.10491 | 278775.5266 | 0.12904 | 0.024128
281612.9259 | 0.1051 | 281612.9259 | 0.13412 | 0.029023
284479.2045 | 0.10186 | 284479.2045 | 0.12595 | 0.024088
287374.6563 | 0.091386 | 287374.6563 | 0.10095 | 0.0095597
290299.5782 | 0.10156 | 290299.5782 | 0.12804 | 0.026482
293254.2703 | 0.097266 | 293254.2703 | 0.125 | 0.027738
296239.0354 | 0.09838 | 296239.0354 | 0.11924 | 0.020857
299254.1797 | 0.083 | 299254.1797 | 0.12179 | 0.038791
302300.0123 | 0.09547 | 302300.0123 | 0.10737 | 0.011899
305376.8457 | 0.095459 | 305376.8457 | 0.12 | 0.024543
308484.9954 | 0.094479 | 308484.9954 | 0.11144 | 0.016958
311624.78 | 0.080874 | 311624.78 | 0.10564 | 0.024764
314796.5216 | 0.091145 | 314796.5216 | 0.11405 | 0.022906
318000.5455 | 0.090755 | 318000.5455 | 0.12148 | 0.030723
321237.1801 | 0.090162 | 321237.1801 | 0.10308 | 0.012915
324506.7575 | 0.082295 | 324506.7575 | 0.10831 | 0.026019
327809.6129 | 0.086593 | 327809.6129 | 0.11469 | 0.028094
331146.085 | 0.086573 | 331146.085 | 0.12423 | 0.037662
334516.516 | 0.081148 | 334516.516 | 0.10483 | 0.023679
337921.2515 | 0.083227 | 337921.2515 | 0.12594 | 0.042708
341360.6407 | 0.080702 | 341360.6407 | 0.11374 | 0.03304
344835.0363 | 0.081849 | 344835.0363 | 0.11564 | 0.033795
348344.7945 | 0.069013 | 348344.7945 | 0.096254 | 0.027241
351890.2753 | 0.081381 | 351890.2753 | 0.13004 | 0.048654
355471.8423 | 0.078496 | 355471.8423 | 0.12514 | 0.046647
359089.8628 | 0.070762 | 359089.8628 | 0.11298 | 0.042216
362744.7078 | 0.075789 | 362744.7078 | 0.10944 | 0.033652
366436.752 | 0.073341 | 366436.752 | 0.115 | 0.041658
370166.3742 | 0.073993 | 370166.3742 | 0.11815 | 0.044155
373933.9568 | 0.07114 | 373933.9568 | 0.10741 | 0.036268
377739.8861 | 0.073603 | 377739.8861 | 0.13719 | 0.063586
381584.5525 | 0.069255 | 381584.5525 | 0.12318 | 0.053927
385468.3502 | 0.06408 | 385468.3502 | 0.10807 | 0.043994
389391.6775 | 0.068599 | 389391.6775 | 0.11567 | 0.047069
393354.9368 | 0.069537 | 393354.9368 | 0.12154 | 0.052007
397358.5344 | 0.056556 | 397358.5344 | 0.11871 | 0.062155
401402.8809 | 0.065262 | 401402.8809 | 0.096559 | 0.031296
405488.3911 | 0.065785 | 405488.3911 | 0.11181 | 0.046029
409615.484 | 0.054731 | 409615.484 | 0.085293 | 0.030562
413784.5827 | 0.062727 | 413784.5827 | 0.10078 | 0.038051
417996.1149 | 0.062409 | 417996.1149 | 0.10605 | 0.043645
422250.5123 | 0.054409 | 422250.5123 | 0.098019 | 0.04361
426548.2114 | 0.058745 | 426548.2114 | 0.10658 | 0.047833
430889.6527 | 0.058156 | 430889.6527 | 0.099299 | 0.041143
435275.2816 | 0.055785 | 435275.2816 | 0.084917 | 0.029131
439705.5478 | 0.055388 | 439705.5478 | 0.09683 | 0.041442
444180.9056 | 0.052118 | 444180.9056 | 0.11479 | 0.062667
448701.8139 | 0.054424 | 448701.8139 | 0.10391 | 0.049489
453268.7364 | 0.052376 | 453268.7364 | 0.10986 | 0.057483
457882.1414 | 0.03784 | 457882.1414 | 0.097685 | 0.059845
462542.502 | 0.050376 | 462542.502 | 0.11317 | 0.062797
467250.2961 | 0.048816 | 467250.2961 | 0.097646 | 0.04883
472006.0065 | 0.047601 | 472006.0065 | 0.098184 | 0.050583
476810.1209 | 0.047878 | 476810.1209 | 0.10483 | 0.056953
481663.132 | 0.034904 | 481663.132 | 0.10126 | 0.066355
486565.5373 | 0.045135 | 486565.5373 | 0.11289 | 0.067755
491517.8398 | 0.043517 | 491517.8398 | 0.095436 | 0.05192
496520.5472 | 0.04128 | 496520.5472 | 0.099664 | 0.058384
501574.1725 | 0.042302 | 501574.1725 | 0.1214 | 0.079095
506679.234 | 0.026448 | 506679.234 | 0.07476 | 0.048312
511836.2553 | 0.039436 | 511836.2553 | 0.093821 | 0.054384
517045.7651 | 0.036995 | 517045.7651 | 0.078414 | 0.041419
522308.2977 | 0.038993 | 522308.2977 | 0.10212 | 0.063122
527624.3928 | 0.037851 | 527624.3928 | 0.094321 | 0.05647
532994.5955 | 0.033119 | 532994.5955 | 0.10001 | 0.066893
538419.4566 | 0.034617 | 538419.4566 | 0.08566 | 0.051043
543899.5323 | 0.024781 | 543899.5323 | 0.096989 | 0.072208
549435.3847 | 0.032795 | 549435.3847 | 0.087022 | 0.054227
555027.5814 | 0.021106 | 555027.5814 | 0.083524 | 0.062418
560676.696 | 0.031779 | 560676.696 | 0.08863 | 0.056851
566383.3077 | 0.025223 | 566383.3077 | 0.092613 | 0.06739
572148.0017 | 0.029361 | 572148.0017 | 0.086997 | 0.057636
577971.3692 | 0.026176 | 577971.3692 | 0.07837 | 0.052194
583854.0074 | 0.029846 | 583854.0074 | 0.095335 | 0.065489
589796.5196 | 0.025998 | 589796.5196 | 0.094603 | 0.068605
595799.5152 | 0.024233 | 595799.5152 | 0.084343 | 0.06011
601863.6097 | 0.025605 | 601863.6097 | 0.087488 | 0.061884
607989.425 | 0.025157 | 607989.425 | 0.084961 | 0.059804
614177.5894 | 0.021581 | 614177.5894 | 0.084016 | 0.062436
620428.7374 | 0.023815 | 620428.7374 | 0.078996 | 0.055182
626743.51 | 0.019082 | 626743.51 | 0.087568 | 0.068486
633122.555 | 0.022604 | 633122.555 | 0.077549 | 0.054944
639566.5263 | 0.014657 | 639566.5263 | 0.073922 | 0.059266
646076.0849 | 0.01914 | 646076.0849 | 0.074604 | 0.055464
652651.8983 | 0.0068067 | 652651.8983 | 0.078437 | 0.071631
659294.6409 | 0.016382 | 659294.6409 | 0.072575 | 0.056193
666004.9938 | 0.0087007 | 666004.9938 | 0.059955 | 0.051255
672783.6452 | 0.014925 | 672783.6452 | 0.083157 | 0.068232
679631.2902 | 0.01378 | 679631.2902 | 0.080377 | 0.066596
686548.6312 | 0.016129 | 686548.6312 | 0.07158 | 0.055452
693536.3773 | 0.013989 | 693536.3773 | 0.089513 | 0.075524
700595.2453 | 0.0074157 | 700595.2453 | 0.069514 | 0.062098
707725.9591 | 0.01225 | 707725.9591 | 0.066389 | 0.054139
714929.2497 | 0.003145 | 714929.2497 | 0.067088 | 0.063943
722205.8561 | 0.0091672 | 722205.8561 | 0.06467 | 0.055503
729556.5243 | 0.010946 | 729556.5243 | 0.090525 | 0.079579
736982.0082 | 0.0042672 | 736982.0082 | 0.071646 | 0.067378
744483.0693 | 0.007602 | 744483.0693 | 0.095225 | 0.087623
752060.4767 | 0.0013166 | 752060.4767 | 0.072498 | 0.071181
759715.0077 | 0.0068767 | 759715.0077 | 0.074992 | 0.068115
767447.447 | 0.005269 | 767447.447 | 0.074263 | 0.068994
775258.5877 | -0.0097994 | 775258.5877 | 0.048034 | 0.057833
783149.2308 | 0.0035191 | 783149.2308 | 0.056362 | 0.052842
791120.1855 | 0.0015077 | 791120.1855 | 0.067292 | 0.065785
799172.2692 | -0.0096401 | 799172.2692 | 0.048308 | 0.057948
807306.3077 | -0.0002566 | 807306.3077 | 0.062322 | 0.062578
815523.135 | 5.657e-05 | 815523.135 | 0.057512 | 0.057456
823823.5938 | -0.019458 | 823823.5938 | 0.052465 | 0.071923
832208.5354 | -0.0015053 | 832208.5354 | 0.067164 | 0.068669
840678.8196 | -0.002366 | 840678.8196 | 0.06542 | 0.067786
849235.315 | -0.010191 | 849235.315 | 0.047148 | 0.057339
857878.8991 | -0.0042085 | 857878.8991 | 0.073165 | 0.077373
866610.4582 | -0.0046175 | 866610.4582 | 0.059241 | 0.063858
875430.8879 | -0.0057717 | 875430.8879 | 0.053244 | 0.059016
884341.0926 | -0.015096 | 884341.0926 | 0.051228 | 0.066325
893341.9861 | -0.0069716 | 893341.9861 | 0.057716 | 0.064688
902434.4913 | -0.0092182 | 902434.4913 | 0.057024 | 0.066242
911619.5408 | -0.010615 | 911619.5408 | 0.051815 | 0.06243
920898.0765 | -0.018374 | 920898.0765 | 0.051696 | 0.070071
930271.0498 | -0.010396 | 930271.0498 | 0.051143 | 0.061539
939739.422 | -0.011343 | 939739.422 | 0.049746 | 0.061089
949304.164 | -0.011935 | 949304.164 | 0.061837 | 0.073772
958966.2568 | -0.02168 | 958966.2568 | 0.042468 | 0.064148
968726.691 | -0.016647 | 968726.691 | 0.043497 | 0.060144
978586.4678 | -0.013756 | 978586.4678 | 0.047326 | 0.061082
988546.5981 | -0.016736 | 988546.5981 | 0.060739 | 0.077476
998608.1034 | -0.017691 | 998608.1034 | 0.052469 | 0.070161
1008772.0155 | -0.02098 | 1008772.0155 | 0.045111 | 0.066091
1019039.3768 | -0.038812 | 1019039.3768 | 0.021554 | 0.060366
1029411.24 | -0.020434 | 1029411.24 | 0.053941 | 0.074375
1039888.6689 | -0.020747 | 1039888.6689 | 0.046503 | 0.06725
1050472.7378 | -0.021952 | 1050472.7378 | 0.053971 | 0.075924
1061164.5323 | -0.023082 | 1061164.5323 | 0.044417 | 0.067499
1071965.1487 | -0.024236 | 1071965.1487 | 0.039012 | 0.063248
1082875.6946 | -0.024822 | 1082875.6946 | 0.037788 | 0.06261
1093897.2889 | -0.030027 | 1093897.2889 | 0.029757 | 0.059784
1105031.0619 | -0.038485 | 1105031.0619 | 0.022981 | 0.061465
1116278.1553 | -0.0365 | 1116278.1553 | 0.02099 | 0.05749
1127639.7225 | -0.031684 | 1127639.7225 | 0.03313 | 0.064815
1139116.9286 | -0.030167 | 1139116.9286 | 0.031246 | 0.061413
1150710.9506 | -0.03068 | 1150710.9506 | 0.03891 | 0.06959
1162422.9775 | -0.030926 | 1162422.9775 | 0.037821 | 0.068747
1174254.2104 | -0.031001 | 1174254.2104 | 0.034925 | 0.065926
1186205.8624 | -0.034306 | 1186205.8624 | 0.020822 | 0.055128
1198279.1594 | -0.034691 | 1198279.1594 | 0.023153 | 0.057845
1210475.3392 | -0.035233 | 1210475.3392 | 0.024975 | 0.060208
1222795.6528 | -0.036138 | 1222795.6528 | 0.035632 | 0.07177
1235241.3635 | -0.038246 | 1235241.3635 | 0.0245 | 0.062746
1247813.7476 | -0.038783 | 1247813.7476 | 0.028039 | 0.066822
1260514.0945 | -0.040044 | 1260514.0945 | 0.030859 | 0.070903
1273343.7064 | -0.042801 | 1273343.7064 | 0.020607 | 0.063408
1286303.8992 | -0.046493 | 1286303.8992 | 0.013677 | 0.060169
1299396.0019 | -0.051993 | 1299396.0019 | 0.0042079 | 0.056201
1312621.357 | -0.056702 | 1312621.357 | -0.005482 | 0.05122
1325981.3208 | -0.050313 | 1325981.3208 | 0.011465 | 0.061778
1339477.2633 | -0.047853 | 1339477.2633 | 0.018588 | 0.066441
1353110.5687 | -0.049055 | 1353110.5687 | -0.0064852 | 0.042569
1366882.6349 | -0.050318 | 1366882.6349 | 0.0057134 | 0.056032
1380794.8742 | -0.051172 | 1380794.8742 | 0.016903 | 0.068075
1394848.7135 | -0.052047 | 1394848.7135 | 0.0052461 | 0.057293
1409045.5938 | -0.061811 | 1409045.5938 | 0.02258 | 0.084392
1423386.9711 | -0.060183 | 1423386.9711 | 0.0173 | 0.077483
1437874.3161 | -0.05695 | 1437874.3161 | 0.0030958 | 0.060045
1452509.1144 | -0.058127 | 1452509.1144 | 0.011378 | 0.069504
1467292.8668 | -0.059083 | 1467292.8668 | -0.0035977 | 0.055485
1482227.0895 | -0.067821 | 1482227.0895 | -0.0078154 | 0.060005
1497313.3138 | -0.064437 | 1497313.3138 | -0.016018 | 0.048419
1512553.0869 | -0.06424 | 1512553.0869 | -0.0059409 | 0.058299
1527947.9717 | -0.065087 | 1527947.9717 | -0.0072612 | 0.057825
1543499.5468 | -0.075196 | 1543499.5468 | -0.020501 | 0.054694
1559209.407 | -0.068478 | 1559209.407 | -0.014946 | 0.053531
1575079.1634 | -0.070879 | 1575079.1634 | -0.014249 | 0.05663
1591110.4435 | -0.07341 | 1591110.4435 | -0.016249 | 0.057161
1607304.8912 | -0.074573 | 1607304.8912 | -0.013472 | 0.061101
1623664.1673 | -0.074892 | 1623664.1673 | -0.015765 | 0.059127
1640189.9494 | -0.079343 | 1640189.9494 | -0.026254 | 0.053089
1656883.9322 | -0.07871 | 1656883.9322 | -0.019735 | 0.058975
1673747.8277 | -0.079677 | 1673747.8277 | -0.023981 | 0.055695
1690783.3652 | -0.095516 | 1690783.3652 | -0.032154 | 0.063362
1707992.2918 | -0.082483 | 1707992.2918 | -0.025487 | 0.056996
1725376.3722 | -0.086879 | 1725376.3722 | -0.03816 | 0.048719
1742937.3891 | -0.08628 | 1742937.3891 | -0.033426 | 0.052854
1760677.1435 | -0.088083 | 1760677.1435 | -0.033409 | 0.054674
1778597.4544 | -0.090488 | 1778597.4544 | -0.03638 | 0.054107
1796700.1597 | -0.09248 | 1796700.1597 | -0.039081 | 0.053399
1814987.1158 | -0.095348 | 1814987.1158 | -0.042589 | 0.05276
1833460.198 | -0.096103 | 1833460.198 | -0.038047 | 0.058056
1852121.3006 | -0.098576 | 1852121.3006 | -0.042968 | 0.055608
1870972.3375 | -0.10135 | 1870972.3375 | -0.052623 | 0.048726
1890015.2416 | -0.10237 | 1890015.2416 | -0.051655 | 0.050712
1909251.966 | -0.11001 | 1909251.966 | -0.059062 | 0.050948
1928684.4832 | -0.10723 | 1928684.4832 | -0.054033 | 0.053194
1948314.7861 | -0.11376 | 1948314.7861 | -0.060852 | 0.052907
1968144.8879 | -0.11062 | 1968144.8879 | -0.065085 | 0.045535
1988176.8219 | -0.11379 | 1988176.8219 | -0.070713 | 0.04308
2008412.6426 | -0.12796 | 2008412.6426 | -0.081021 | 0.046937
2028854.425 | -0.118 | 2028854.425 | -0.067498 | 0.050498
2049504.2655 | -0.12143 | 2049504.2655 | -0.06852 | 0.052905
2070364.2818 | -0.12806 | 2070364.2818 | -0.085976 | 0.042089
2091436.6129 | -0.12741 | 2091436.6129 | -0.08047 | 0.046945
2112723.4198 | -0.12832 | 2112723.4198 | -0.08136 | 0.046963
2134226.8856 | -0.13074 | 2134226.8856 | -0.091589 | 0.039152
2155949.2153 | -0.13756 | 2155949.2153 | -0.096215 | 0.041341
2177892.6367 | -0.14222 | 2177892.6367 | -0.10271 | 0.039508
2200059.3999 | -0.13873 | 2200059.3999 | -0.09964 | 0.039092
2222451.7781 | -0.14242 | 2222451.7781 | -0.1032 | 0.039218
2245072.0678 | -0.14571 | 2245072.0678 | -0.1101 | 0.035608
2267922.5886 | -0.14802 | 2267922.5886 | -0.10444 | 0.043587
2291005.6837 | -0.15227 | 2291005.6837 | -0.11814 | 0.034136
2314323.7204 | -0.16046 | 2314323.7204 | -0.11832 | 0.042142
2337879.09 | -0.17128 | 2337879.09 | -0.13735 | 0.033928
2361674.2079 | -0.18266 | 2361674.2079 | -0.14434 | 0.038321
2385711.5145 | -0.17261 | 2385711.5145 | -0.13668 | 0.035933
2409993.4746 | -0.17489 | 2409993.4746 | -0.14283 | 0.032066
2434522.5785 | -0.17926 | 2434522.5785 | -0.15074 | 0.028527
2459301.3415 | -0.19621 | 2459301.3415 | -0.16755 | 0.028664
2484332.3047 | -0.19155 | 2484332.3047 | -0.15772 | 0.033831
2509618.0351 | -0.18754 | 2509618.0351 | -0.16178 | 0.025761
2535161.1255 | -0.18667 | 2535161.1255 | -0.16437 | 0.022294
2560964.1956 | -0.18899 | 2560964.1956 | -0.16421 | 0.024781
2587029.8914 | -0.19313 | 2587029.8914 | -0.16689 | 0.026244
2613360.8859 | -0.197 | 2613360.8859 | -0.17648 | 0.020515
2639959.8793 | -0.20075 | 2639959.8793 | -0.18047 | 0.020279
2666829.5994 | -0.22703 | 2666829.5994 | -0.20733 | 0.0197
2693972.8016 | -0.20996 | 2693972.8016 | -0.19153 | 0.018429
2721392.2695 | -0.21369 | 2721392.2695 | -0.19908 | 0.014609
2749090.8148 | -0.2175 | 2749090.8148 | -0.22083 | -0.003326
2777071.2782 | -0.23861 | 2777071.2782 | -0.22479 | 0.013816
2805336.5289 | -0.22629 | 2805336.5289 | -0.20804 | 0.01825
2833889.4656 | -0.23107 | 2833889.4656 | -0.2158 | 0.01527
2862733.0163 | -0.24964 | 2862733.0163 | -0.23971 | 0.009932
2891870.139 | -0.24089 | 2891870.139 | -0.23493 | 0.0059638
2921303.8216 | -0.24649 | 2921303.8216 | -0.23831 | 0.0081816
2951037.0825 | -0.25071 | 2951037.0825 | -0.24478 | 0.0059353
2981072.9709 | -0.25554 | 2981072.9709 | -0.24725 | 0.0082865
3011414.5669 | -0.26156 | 3011414.5669 | -0.25786 | 0.0037002
3042064.9822 | -0.26566 | 3042064.9822 | -0.26209 | 0.0035674
3073027.3597 | -0.27157 | 3073027.3597 | -0.27197 | -0.00039522
3104304.8748 | -0.27739 | 3104304.8748 | -0.27969 | -0.0023057
3135900.7349 | -0.28243 | 3135900.7349 | -0.2888 | -0.0063612
3167818.1802 | -0.30312 | 3167818.1802 | -0.30941 | -0.006298
3200060.4837 | -0.29421 | 3200060.4837 | -0.29834 | -0.0041317
3232630.952 | -0.29982 | 3232630.952 | -0.30514 | -0.0053185
3265532.925 | -0.32017 | 3265532.925 | -0.33069 | -0.010517
3298769.7769 | -0.31201 | 3298769.7769 | -0.32589 | -0.013876
3332344.9162 | -0.31864 | 3332344.9162 | -0.3382 | -0.019553
3366261.7858 | -0.32457 | 3366261.7858 | -0.34049 | -0.01592
3400523.864 | -0.33849 | 3400523.864 | -0.35458 | -0.016093
3435134.6644 | -0.34295 | 3435134.6644 | -0.36342 | -0.020467
3470097.7362 | -0.34452 | 3470097.7362 | -0.36919 | -0.024666
3505416.6649 | -0.35106 | 3505416.6649 | -0.37904 | -0.027977
3541095.0725 | -0.35874 | 3541095.0725 | -0.38489 | -0.026144
3577136.6177 | -0.36606 | 3577136.6177 | -0.39271 | -0.026656
3613544.9966 | -0.37351 | 3613544.9966 | -0.40444 | -0.030931
3650323.9429 | -0.38072 | 3650323.9429 | -0.41428 | -0.033556
3687477.2282 | -0.38824 | 3687477.2282 | -0.42688 | -0.038648
3725008.6625 | -0.39557 | 3725008.6625 | -0.43725 | -0.041683
3762922.0948 | -0.40317 | 3762922.0948 | -0.44734 | -0.044173
3801221.413 | -0.41637 | 3801221.413 | -0.46507 | -0.048691
3839910.5446 | -0.42742 | 3839910.5446 | -0.47584 | -0.048421
3878993.4573 | -0.4274 | 3878993.4573 | -0.4795 | -0.052102
3918474.159 | -0.43658 | 3918474.159 | -0.49353 | -0.056951
3958356.6983 | -0.44473 | 3958356.6983 | -0.50509 | -0.060357
3998645.1653 | -0.46106 | 3998645.1653 | -0.52667 | -0.065607
4039343.6915 | -0.46295 | 4039343.6915 | -0.52973 | -0.066776
4080456.4505 | -0.47173 | 4080456.4505 | -0.54243 | -0.070703
4121987.6584 | -0.48193 | 4121987.6584 | -0.55259 | -0.070664
4163941.5742 | -0.49029 | 4163941.5742 | -0.55711 | -0.06682
4206322.5003 | -0.50577 | 4206322.5003 | -0.58748 | -0.081711
4249134.7828 | -0.50937 | 4249134.7828 | -0.59626 | -0.086898
4292382.8121 | -0.52241 | 4292382.8121 | -0.6043 | -0.081893
4336071.0232 | -0.52912 | 4336071.0232 | -0.61907 | -0.089958
4380203.8965 | -0.5393 | 4380203.8965 | -0.63394 | -0.094644
4424785.9575 | -0.56951 | 4424785.9575 | -0.66618 | -0.096671
4469821.7783 | -0.56114 | 4469821.7783 | -0.66328 | -0.10214
4515315.9773 | -0.57166 | 4515315.9773 | -0.68213 | -0.11047
4561273.2198 | -0.58469 | 4561273.2198 | -0.695 | -0.11031
4607698.2187 | -0.61702 | 4607698.2187 | -0.73065 | -0.11363
4654595.735 | -0.60604 | 4654595.735 | -0.72535 | -0.1193
4701970.5779 | -0.617 | 4701970.5779 | -0.73962 | -0.12262
4749827.6056 | -0.62914 | 4749827.6056 | -0.76129 | -0.13215
4798171.726 | -0.64159 | 4798171.726 | -0.77551 | -0.13392
4847007.8966 | -0.6538 | 4847007.8966 | -0.79192 | -0.13811
4896341.1257 | -0.66637 | 4896341.1257 | -0.81004 | -0.14367
4946176.4722 | -0.67863 | 4946176.4722 | -0.8271 | -0.14847
4996519.0469 | -0.6926 | 4996519.0469 | -0.84408 | -0.15148
5047374.0122 | -0.72354 | 5047374.0122 | -0.88111 | -0.15757
5098746.5834 | -0.71846 | 5098746.5834 | -0.88227 | -0.16381
5150642.0287 | -0.73132 | 5150642.0287 | -0.8989 | -0.16758
5203065.6699 | -0.744 | 5203065.6699 | -0.92022 | -0.17622
5256022.8831 | -0.76464 | 5256022.8831 | -0.94378 | -0.17913
5309519.0989 | -0.77126 | 5309519.0989 | -0.95529 | -0.18403
5363559.8035 | -0.78811 | 5363559.8035 | -0.97824 | -0.19013
5418150.5386 | -0.79819 | 5418150.5386 | -0.995 | -0.19681
5473296.9026 | -0.81471 | 5473296.9026 | -1.0155 | -0.20077
5529004.5506 | -0.82472 | 5529004.5506 | -1.0343 | -0.20957
5585279.1954 | -0.84765 | 5585279.1954 | -1.0604 | -0.21273
5642126.608 | -0.85096 | 5642126.608 | -1.0718 | -0.22081
5699552.6181 | -0.86483 | 5699552.6181 | -1.0864 | -0.22156
5757563.1146 | -0.87947 | 5757563.1146 | -1.1096 | -0.23017
5816164.0466 | -0.91422 | 5816164.0466 | -1.1516 | -0.2374
5875361.4234 | -0.90428 | 5875361.4234 | -1.1443 | -0.23999
5935161.3158 | -0.91627 | 5935161.3158 | -1.1661 | -0.24982
5995569.8563 | -0.92918 | 5995569.8563 | -1.1836 | -0.25445
6056593.2397 | -0.94196 | 6056593.2397 | -1.2112 | -0.2692
6118237.7238 | -0.95434 | 6118237.7238 | -1.2281 | -0.27376
6180509.6304 | -0.96805 | 6180509.6304 | -1.2471 | -0.27901
6243415.3454 | -1.0034 | 6243415.3454 | -1.2942 | -0.29082
6306961.3197 | -0.99508 | 6306961.3197 | -1.2873 | -0.29217
6371154.0699 | -1.0066 | 6371154.0699 | -1.3141 | -0.30754
6436000.1791 | -1.0191 | 6436000.1791 | -1.3346 | -0.31552
6501506.297 | -1.0375 | 6501506.297 | -1.3627 | -0.32525
6567679.1414 | -1.0469 | 6567679.1414 | -1.3789 | -0.33199
6634525.4982 | -1.0762 | 6634525.4982 | -1.4167 | -0.34048
6702052.2225 | -1.0747 | 6702052.2225 | -1.4273 | -0.35268
6770266.2392 | -1.0921 | 6770266.2392 | -1.4501 | -0.35799
6839174.5435 | -1.1026 | 6839174.5435 | -1.4715 | -0.36894
6908784.2019 | -1.1178 | 6908784.2019 | -1.4963 | -0.37851
6979102.353 | -1.1328 | 6979102.353 | -1.5289 | -0.39604
7050136.2078 | -1.158 | 7050136.2078 | -1.5605 | -0.4025
7121893.0507 | -1.1953 | 7121893.0507 | -1.6028 | -0.40749
7194380.2404 | -1.2292 | 7194380.2404 | -1.6517 | -0.4225
7267605.2105 | -1.247 | 7267605.2105 | -1.684 | -0.43699
7341575.4701 | -1.2476 | 7341575.4701 | -1.69 | -0.44244
7416298.6049 | -1.2444 | 7416298.6049 | -1.6981 | -0.45368
7491782.2776 | -1.2528 | 7491782.2776 | -1.7166 | -0.46379
7568034.2291 | -1.272 | 7568034.2291 | -1.7496 | -0.47761
7645062.279 | -1.2909 | 7645062.279 | -1.777 | -0.48611
7722874.3265 | -1.3118 | 7722874.3265 | -1.8118 | -0.49996
7801478.3511 | -1.3298 | 7801478.3511 | -1.8438 | -0.51403
7880882.4137 | -1.355 | 7880882.4137 | -1.877 | -0.52203
7961094.6571 | -1.3699 | 7961094.6571 | -1.9057 | -0.53582
8042123.307 | -1.3908 | 8042123.307 | -1.9436 | -0.55283
8123976.673 | -1.4262 | 8123976.673 | -1.9886 | -0.56234
8206663.1489 | -1.4343 | 8206663.1489 | -2.0071 | -0.57283
8290191.2144 | -1.456 | 8290191.2144 | -2.0385 | -0.5825
8374569.4351 | -1.4788 | 8374569.4351 | -2.0774 | -0.59859
8459806.464 | -1.5017 | 8459806.464 | -2.1178 | -0.61615
8545911.0421 | -1.5251 | 8545911.0421 | -2.1622 | -0.63705
8632891.9993 | -1.5491 | 8632891.9993 | -2.1961 | -0.64703
8720758.2557 | -1.6094 | 8720758.2557 | -2.2737 | -0.66428
8809518.8217 | -1.5994 | 8809518.8217 | -2.2856 | -0.68628
8899182.7998 | -1.6239 | 8899182.7998 | -2.3232 | -0.69928
8989759.3849 | -1.6765 | 8989759.3849 | -2.3996 | -0.72305
9081257.8657 | -1.6768 | 9081257.8657 | -2.419 | -0.74228
9173687.6252 | -1.7183 | 9173687.6252 | -2.4757 | -0.75737
9267058.1421 | -1.732 | 9267058.1421 | -2.51 | -0.77798
9361378.9915 | -1.7609 | 9361378.9915 | -2.5641 | -0.80323
9456659.846 | -1.7909 | 9456659.846 | -2.6153 | -0.82436
9552910.4765 | -1.8282 | 9552910.4765 | -2.6756 | -0.84733
9650140.7536 | -1.8633 | 9650140.7536 | -2.7336 | -0.87035
9748360.648 | -1.8918 | 9748360.648 | -2.7817 | -0.88991
9847580.2324 | -1.9199 | 9847580.2324 | -2.8327 | -0.91281
9947809.6814 | -1.9551 | 9947809.6814 | -2.8822 | -0.92708
10049059.2737 | -1.9911 | 10049059.2737 | -2.9364 | -0.94535
10151339.3924 | -2.0347 | 10151339.3924 | -3.0041 | -0.96942
10254660.5262 | -2.0689 | 10254660.5262 | -3.0628 | -0.99391
10359033.2706 | -2.113 | 10359033.2706 | -3.115 | -1.0019
10464468.3291 | -2.1511 | 10464468.3291 | -3.1768 | -1.0257
10570976.514 | -2.1945 | 10570976.514 | -3.2437 | -1.0492
10678568.7475 | -2.2381 | 10678568.7475 | -3.3032 | -1.0651
10787256.0633 | -2.2841 | 10787256.0633 | -3.3747 | -1.0906
10897049.6072 | -2.3302 | 10897049.6072 | -3.4418 | -1.1117
11007960.6384 | -2.3766 | 11007960.6384 | -3.5037 | -1.1271
11120000.5308 | -2.4256 | 11120000.5308 | -3.5749 | -1.1492
11233180.7742 | -2.4742 | 11233180.7742 | -3.6449 | -1.1707
11347512.975 | -2.5309 | 11347512.975 | -3.7232 | -1.1922
11463008.8579 | -2.5775 | 11463008.8579 | -3.794 | -1.2165
11579680.2671 | -2.6341 | 11579680.2671 | -3.8727 | -1.2385
11697539.1671 | -2.6833 | 11697539.1671 | -3.9438 | -1.2605
11816597.6442 | -2.7393 | 11816597.6442 | -4.0311 | -1.2918
11936867.908 | -2.7955 | 11936867.908 | -4.1171 | -1.3216
12058362.292 | -2.8541 | 12058362.292 | -4.1991 | -1.345
12181093.2554 | -2.9165 | 12181093.2554 | -4.2934 | -1.3769
12305073.3842 | -2.9814 | 12305073.3842 | -4.3877 | -1.4062
12430315.3926 | -3.1254 | 12430315.3926 | -4.5599 | -1.4345
12556832.1241 | -3.1229 | 12556832.1241 | -4.5815 | -1.4586
12684636.5528 | -3.198 | 12684636.5528 | -4.6875 | -1.4895
12813741.7851 | -3.2821 | 12813741.7851 | -4.7791 | -1.497
12944161.0606 | -3.4132 | 12944161.0606 | -4.9358 | -1.5226
13075907.7539 | -3.4708 | 13075907.7539 | -4.995 | -1.5242
13208995.3753 | -3.578 | 13208995.3753 | -5.1139 | -1.5359
13343437.5731 | -3.6941 | 13343437.5731 | -5.2183 | -1.5242
13479248.1342 | -3.8177 | 13479248.1342 | -5.3242 | -1.5065
13616440.9859 | -3.9711 | 13616440.9859 | -5.4774 | -1.5063
13755030.1973 | -4.0797 | 13755030.1973 | -5.5638 | -1.4841
13895029.9806 | -4.2919 | 13895029.9806 | -5.7631 | -1.4712
14036454.6927 | -4.3383 | 14036454.6927 | -5.7978 | -1.4595
14179318.8368 | -4.4688 | 14179318.8368 | -5.9245 | -1.4557
14323637.0634 | -4.6184 | 14323637.0634 | -6.099 | -1.4806
14469424.1723 | -4.7431 | 14469424.1723 | -6.2562 | -1.5131
14616695.1139 | -4.8216 | 14616695.1139 | -6.3604 | -1.5388
14765464.9907 | -4.8815 | 14765464.9907 | -6.4702 | -1.5887
14915749.0591 | -4.9588 | 14915749.0591 | -6.6141 | -1.6553
15067562.7307 | -5.0899 | 15067562.7307 | -6.8249 | -1.7349
15220921.5737 | -5.0956 | 15220921.5737 | -6.9184 | -1.8227
15375841.3152 | -5.1576 | 15375841.3152 | -7.0772 | -1.9196
15532337.842 | -5.2361 | 15532337.842 | -7.2914 | -2.0553
15690427.2029 | -5.2952 | 15690427.2029 | -7.4665 | -2.1713
15850125.6097 | -5.3233 | 15850125.6097 | -7.6113 | -2.288
16011449.4395 | -5.3556 | 16011449.4395 | -7.7674 | -2.4118
16174415.236 | -5.4077 | 16174415.236 | -7.9698 | -2.5622
16339039.7113 | -5.4736 | 16339039.7113 | -8.1812 | -2.7076
16505339.7475 | -5.5025 | 16505339.7475 | -8.3373 | -2.8348
16673332.3985 | -5.5475 | 16673332.3985 | -8.5537 | -3.0062
16843034.892 | -5.5879 | 16843034.892 | -8.7239 | -3.136
17014464.6309 | -5.6264 | 17014464.6309 | -8.9265 | -3.3001
17187639.1952 | -5.6645 | 17187639.1952 | -9.11 | -3.4454
17362576.3439 | -5.7552 | 17362576.3439 | -9.3232 | -3.568
17539294.0167 | -5.7338 | 17539294.0167 | -9.4269 | -3.6931
17717810.336 | -5.792 | 17717810.336 | -9.6003 | -3.8083
17898143.6084 | -5.9043 | 17898143.6084 | -9.8448 | -3.9405
18080312.3271 | -5.8945 | 18080312.3271 | -10.0509 | -4.1563
18264335.1734 | -5.807 | 18264335.1734 | -10.1659 | -4.3589
18450231.0188 | -5.8657 | 18450231.0188 | -10.1973 | -4.3316
18638018.9268 | -5.7174 | 18638018.9268 | -10.3448 | -4.6274
18827718.1549 | -5.7221 | 18827718.1549 | -10.5781 | -4.856
19019348.1567 | -5.6784 | 19019348.1567 | -10.7462 | -5.0678
19212928.5838 | -5.7105 | 19212928.5838 | -10.9704 | -5.2599
19408479.2878 | -5.7627 | 19408479.2878 | -11.1712 | -5.4085
19606020.3223 | -6.0057 | 19606020.3223 | -11.6535 | -5.6478
19805571.9451 | -5.9039 | 19805571.9451 | -11.6857 | -5.7818
20007154.62 | -5.9438 | 20007154.62 | -11.7793 | -5.8355
20210789.0193 | -6.0028 | 20210789.0193 | -12.0233 | -6.0206
20416496.0256 | -6.0794 | 20416496.0256 | -12.1204 | -6.041
20624296.734 | -6.1255 | 20624296.734 | -12.2116 | -6.0861
20834212.4545 | -6.1404 | 20834212.4545 | -12.3475 | -6.2071
21046264.7137 | -6.3557 | 21046264.7137 | -12.6628 | -6.307
21260475.2576 | -6.1809 | 21260475.2576 | -12.6788 | -6.4979
21476866.0533 | -6.2384 | 21476866.0533 | -12.8993 | -6.6609
21695459.2916 | -6.2906 | 21695459.2916 | -13.1043 | -6.8136
21916277.3891 | -6.2807 | 21916277.3891 | -13.1847 | -6.904
22139342.9907 | -6.2674 | 22139342.9907 | -13.4043 | -7.1369
22364678.9715 | -6.4826 | 22364678.9715 | -13.9134 | -7.4308
22592308.4398 | -6.3118 | 22592308.4398 | -13.8776 | -7.5657
22822254.7388 | -6.3352 | 22822254.7388 | -14.13 | -7.7947
23054541.4493 | -6.3525 | 23054541.4493 | -14.3541 | -8.0016
23289192.3924 | -6.3651 | 23289192.3924 | -14.5512 | -8.1861
23526231.6312 | -6.4507 | 23526231.6312 | -14.5174 | -8.0668
23765683.4742 | -6.3517 | 23765683.4742 | -14.8623 | -8.5106
24007572.4769 | -6.3402 | 24007572.4769 | -15.0008 | -8.6606
24251923.445 | -6.3289 | 24251923.445 | -14.9581 | -8.6292
24498761.4365 | -6.3124 | 24498761.4365 | -14.986 | -8.6736
24748111.7646 | -6.3425 | 24748111.7646 | -14.9884 | -8.6459
|  | 25000000 | -16.1917